## Slide 1: Performance 3 point Brenner

## Slide 2: Focal Plane differences
DAPI: 48.1um
A488 ezrin:52.2um
A647 Na/K ATPase: 54.5um
DAPI: 31.1um
488 auto-fl: 26um

## Slide 3: DAPI FWHM (cropped)
17um
Focus: 47.5um
FWHM: 12um
Focus: 46.2um
FWHM: 12um
Focus: 46.0 um
FWHM: 9 um

## Slide 4: DAPI FWHM (cropped)
17um
18um
12um
12um

## Slide 5: Low S/N

## Slide 6: FWHM Vs. Binning
Why FWHM? FWHM is a defining property in scan range of 3PAF (3 point autofocus). If 3 points are gathered under FWHM w/ at least one on either side of focus, the result will be accurate.
1
2
4
8
16
32
Chart: Bin Vs FWHM
| 1 | 2 | 4 | 8 | 16 | 32
 | 11.3 | 13.6 | 16.6 | 24 | 33.3 | 50
Large FWHM up from increase in binning
Center of parabola is roughly the same

## Slide 7: 3 Point Search Solution Vs Binning
Point 1 in first 3rd
Search Range = FWHM at 32 bin level. Each of 3 points is chosen from jointing 3rd sections
Point 2 in Second 3rd
Point 3 in Last 3rd
Focus is contained within this third region somewhere
Possible solutions for 32 bin
Possible solutions for 8 bin
Possible solutions for 2 bin
Possible solutions for 16 bin
Best solutions = 96, 97, 99

## Slide 8: S/N and binning

## Slide 9: What about Stains for smaller distributions?
Ezrin is a finer stained feature than DAPI
Binning effectiveness may be balance between S/N and focused feature size
DAPI image was best at 32 bin level.
Ezrin does not appear to be. Best is 8 bin level.
Chart: Bin Vs ezrin Score range
| 1 | 2 | 4 | 8 | 16 | 32
 | 0.6 | 2 | 5.5 | 11 | 10.8 | 7

## Slide 10: What about not much area occupied?
Binning certainly helps images with low area occupation
At first its tempting to say it will bin away to nothing and hurt it.
Possibly it helps by effectively increasing its weight in the overall image on a pixel by pixel basis.
Easy to see that best is bin 32. While it is noisier, it is max FWHM while maintaining sufficiently low noise levels.

## Slide 11: Can I focus a single Nucleus?
Bin=1
Bin=8
Bin=16
Bin=32
It isn’t perfect, but a single nuclei is focusable!
Noise problems would be a concern with the 3 point method
Red bar is at known focal plane location

## Slide 12: What about auto bin selection?
Its clear no exact bin level is best for all structures and areas occupied.
What if I selected the bin level based on max range of the 3 points I obtained?

## Slide 13: Lets test the new engine out!
64um
80um
124um
Output solution!
100um

## Slide 14: New bin method

## Slide 15: Skipping bin vs brenner skipping range vs true bin
Brenner skip = 2, skip bin =16
True bin = 1
Brenner skip = 32, skip bin =1
True bin = 1
Brenner skip = 2, skip bin =1
True bin = 16
Brenner skip = 8, skip bin =1
True bin = 4
Brenner skip = 1, skip bin =1
True bin = 32
Brenner skip = 4, skip bin =1
True bin = 8

## Slide 16: Kuwahara Filter
Rapid noise removal filter
Before
After

## Slide 17: Kuwahara Brenner Score Space
w/o
w/
Single nuc
Full DAPI

## Slide 18: Thoughts
‘Wings’ on brenner spectra seem to be noise related.

## Slide 19: Brenner Score “Mask”

## Slide 20
Kuw = on, mask = on, 1 true bin
Kuw = on, mask = on, 4 true bin
Kuw = off, mask = on, 1 true bin
Kuw = off, mask = on, 4 true bin
Kuw = off, mask = off, 4 true bin
Kuw = on, mask = off, 4 true bin

## Slide 21
w/ kuwahara
w/o kuwahara

## Slide 22
Nothing
Median
Butterworth
Median(butterworths)
Kuwahara
No Bin