## Slide 1: July Update

## Slide 2: Flexi-TEER Website and Program
Project selected as part of first cohort of mentees for open-source project development with Open Hardware Makers
Project will be listed in upcoming Open Hardware Development guide
Website developed (still in progress) to distribute device and engage potential contributers

## Slide 3: Transwells? Resistance? TEER?

## Slide 4
*
Lets approximate as:
But why? This is nonsensical. That is a constant value. It cant just go to infinity! The solution lies in the fact that the above solution was never taking the voltage drop over a capacitor into account. It was built assuming all voltage dropped over the media and TJ resistance.
1
1
2
2
What now though? . This would dictate that we knew with 100% certainty what the time of the voltage reading was from the DC current being applied and the dynamics of the current over time, i.e. solve
*dt
Real hard to solve though 

## Slide 5: Modular Breakdown of Flexi-TEER + Schematic
Generate DC square Wave with function generator
Generate AC Wave
Read off voltages via instrumentation amp
Send into Differential Amp
Send into ADC
Amplify Signal

## Slide 6: AC conversion issue
Single Pole Dual Throw switches (SPDT) in an H-Bridge config are great at converting 0-(+V) signal into (–V)-(+V) at same freq
Large amounts of noise are introduced when current is squeezed down.

## Slide 7: Jose’s Proposed Circuit
AC generation + CC
Sample voltage and find difference
Amplify and send to ADC
Better Charge Pump
Isolate reference voltage

## Slide 8: Deriving AC CC unit
If
Then
if
If
This only works for square waves! Sine waves will alter current levels and not just change direction

## Slide 9: CyCIF Multiplex Technique
Multiplex = getting lots and lots of markers on same tissue or sample
Cyclic technique where instead of removing the old fluorophores, we bleach them to get rid of them.
Cycle 1:
Villin
DAPI
Cycle 2:
LAMP1
Phalloidin
63x
LSM 880

## Slide 10: Automation Needs
Auto Focus – Brenner Score
Auto Expose – Threshold on Histogram
Fluidic delivery (large and small scales) – Aquarium and micro stepper peristaltic pump for large and small respectively.
Computational control of microscope and robotic systems – Pycromanager and micromanager
Robotic control system – Arduino and breakout motor drivers
Auto generation of focused surfaces - micromagellan
Auto acquisitions - micromagellan

## Slide 12: New pump system and the “Invisible” coverslip
Issues with Original System
Large stage movement
150-200uL dead volume
Incomplete washings
Diluted stain solution
Limited room for adding more cycles
Large space taken up
New System
Solution storage
Liquid selector (pinch valve)
Liquid Pump (Peristaltic Pump)
Invisi-Slip Chamber

## Slide 13: Auto-CIF Control Board V 3.0
BNC connectors to modulate pinch valve on-off status via TTL waves
On board voltage modulators/PSU
Stepper and DC motor drivers
Arduino Microcontroller

## Slide 14: Flexible Serial Command Interface

## Slide 15: Invisi-Slip System
Glass Slide (Tissue side up)
Objective
FEP Film
100um Gap
1600um
2500um
WD
Fluid Flow
Fluid Flow
Smallish volume on top of entire slide surface (about 400uL)
No dilution of stain solution
FEP film has RI =1.333, PBS = 1.33, So if PBS is above and below the FEP film, it is effectively invisible
FEP film is a physical barrier, but not an optical barrier

## Slide 16: Parametric CAD model
Migrated to the FreeCAD platform from TinkerCAD.
Fully Parametric model to instantly tweak every single feature of the model
More power in features like Fillets and chamfers and symmetries

## Slide 17: Current Status Roadmap for Invisi-Slip
Basic CAD Design
Prototype Chamber
Identify new pump system
Design new PCB to interface Arduino with new System
Build Pump System
Finalize Chamber Design and confirm waterproof even in high pH + H2O2 conditions
Tweak Arduino Code to run new system
We are currently here

## Slide 20: Auto Exposure
Set intensity of the 98% brightest pixel to be at X fraction of dynamic range
Add in sheet of algorithm

## Slide 21: Chicken and Egg problem
Chicken: Brenner based auto focus needs proper exposure to focus correctly
Egg: auto exposure needs to be in focus to determine proper exposure

## Slide 22: Solution
1. Scan over set Z range

## Slide 23: Solution
2. Determine Relative Image Brightness Levels

## Slide 24: Solution
3. Auto Expose on Brightest Z slice

## Slide 25: Solution
4. Run Auto Focus over same range with determined exposure level

## Slide 26: Issues
Hardware
Wrong tubes
Wrong adapter for M32 Objective
Chip Tilt
Reflection
Chamber Leak
Software
Wrong Tiling spacing
Glitch in MM where only one surface can be used for Z points
Outdated Drivers
Computer Crashing
Crashing Camera driver
RAM exhaustion in Acq computer
MM multichannel Acq

## Slide 28
Camera Chip
Sample
Microscope

## Slide 29
Camera Chip
Sample
Microscope

## Slide 30
Camera Chip
Sample
Microscope
What if Stage and Chip are not parallel?

## Slide 31
Camera Chip
Sample
Microscope
What if Stage and Chip are not parallel?

## Slide 32: XY Repeatability
Stage translates beyond small displacements for tiling. It moves stage from under objective to under syringe for staining.
This vector is around an 80mm displacement in X and 5mm in Y.
X axis moved an AVG: 1.14 with std dev of 1.64 per cycle
Y axis moved an AVG: -.29 with a std dev of 0.49 per cycle
Chart: Displacement as Function of Cycle
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
X Axis | 0 | 1 | 4 | 3 | 4 | 4 | 8 | 8
Y Axis | 0 | -1 | -1 | -1 | -1 | -1 | -2 | -2
X is translated way way further than Y so its expected that it is less accurate.
Slide was taped to blackened dog tags for this experiment. The data pointed to this moving on the stage as it had a massive jump of 54 pixels in the Y from image 1 to 2. Its completely possible and likely that the actual stage accuracy is better than this as this is stage plus movement of slide on stage. Will repeat when get a drain chamber that firmly grasps slide to eliminate its movement.

## Slide 33: X and Y Repeatiblity
Here are the 3rd and 10th cycle images overlaid

## Slide 34: Reflection
Top: Image with blackened anodized Al dogtag behind slide. 1000ms exposure
Bottom: Image with 3D printed blue plastic housing behind slide. 2ms exposure
* Experimental Notes.
Top image was taken and then objective raised, Al dogtag was removed and slide reseated into plastic socket, objective lowered and reimaged.
No power cycles occurred. It was in the exact same state.

## Slide 35: Ultimate Reflection Solution
Acktar’s Spectral Black Film
Market leader in maximal absorbing and low reflectance materials

## Slide 37: Bleaching Effectiveness
ECAD-488
ECAD-488 Bleached
Ezrin-488
Ezrin-488 Bleached

## Slide 38
Cycle 1 DAPI
Cycle 1 MUC2
Cycle 1 Lysozyme A
Cycle 2 Ezrin
Cycle 2 Gamma Actin

## Slide 39: Dye Kinetics
Goal = to optimize and document bleaching kinetics of CyCIF compatible dyes
Modulate known variables (pH and % H202)
Find relevant decay kinetics (half life)
Background
Expected decay kinetics form
Why are decays usually exponential? Decays happen due to per molecule occurrences. There is a % chance a molecule will be bleached per unit time. This chance is the same per molecule. Thus, the magnitude of change is proportional to the number of unbleached molecules remaining.
In other words: -> (solve ODE for this solution).

## Slide 40
Chart: A647 Kinetics with additions