Bring up report for Autocyplex.

# Orientation/System Overview

# Basic hardware Testing

## Illumination Profile

## 2.2 Resolution

## 2.3 Reflection

## 2.4 Invisislip vs water imaging comparison

## 2.5 Planar focal plane flatness

## 2.6 chromatic aberration


# Software

## Background Subtraction

Looking at line profile over a Na-K-ATPase stain line, we can see the background subtraction increased peak signal to bckgnd ratio from 1.32 to 2.25.

A screenshot of a computer

Description automatically generated



## Auto Focus

### 12-6-2023 change

Up until now I had been running an autofocus program which entailed taking images in a z stack at every tile and finding the max and centering my acquisition z stack focus map on that. It hit me that this can be simplified and hastened. During the stain phase, I can look through the previous cycles images and recenter the best focus plane for each tile and then update the focus map. This is done during a downtime activity so net time cost on the system is zero. [Presentation\Previous Cycle, Star-Dist Autofocus.pptx](Presentation/Previous%20Cycle,%20Star-Dist%20Autofocus.pptx)

## Auto Exposure

### 2-8-2024 change

Up until now I had been running an autofocus program which entailed taking images in a z stack at every tile and finding the max and centering my acquisition z stack focus map on that. It hit me that this can be simplified and hastened. During the stain phase, I can look through the previous cycles images and recenter the best focus plane for each tile and then update the focus map. This is done during a downtime activity so net time cost on the system is zero. [Presentation\Previous Cycle, Star-Dist Autofocus.pptx](Presentation/Previous%20Cycle,%20Star-Dist%20Autofocus.pptx)

# Issues

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **Problem ID** | **Problem Description** | **Solution** | **Date** | **Person** | **Status** |
| I1 | Core capture got stuck in random spot. No good leads on why. Previously driver in mm had crashed. Not the case here. I restarted image acquisition and it had no issue. For reference, it had captured bleached cycle 1, stain/bleach for cycle 2 and stalled at stained for cycle 3. | Needed updates and a restart. Also must grab pic from core cap or if ask cap again without, will cause it to stall | 11/29/23 | MA | Resolved. |
| F1 | Pressure source loud and unable to sustain pressure >600mBar. | New pump fixed. | 11/28/23 | MA | Resolved |
| F2 | When asking to stop flow, a slow flow still exists. Its very minimal. Seems to make pbs hit the chamber in about 1 hour. | Do new calibration and ask for -3 flow as set point |  |  | Resolved |
| F3 | Only delivered cycle 1 stain in multiplex run. Nothing else was seemingly executed | Works sometimes. Contacted Elveflow to see if maybe instability. | 1/31/2024 | MA | Unknown |
| I2 | 6-2-24 multiplex Cycle 1, A647 was imaged and all pixel values were NaN. | No solution. Might be freak appearance. | 2/6/2024 | MA | Unknown |
|  |  |  |  |  |  |

# Solution Testing

## F1

Pump is mega loud. Record 85db on phone and documents say should be 53db. Performance was not hindered until recently. As of now, it is only able to sustain 600mBar pressure. Solution is to do exchange with Elveflow for new unit as this one is clearly failing. This solution was correct.

## I1

Really hard to say here. It is beyond my abilities to resolve an issue with the core capture stalling out if it is due to architecture. Simply restarting the program made it work though. A possible solution could be to make the capture function wait a bit. If it exceeds a certain time frame, it simply asks it to do it again. Python IDE gave no errors back in this case. It just didn’t proceed on in the code.

Other possible solutions could be to alternative code to capture image. No need to use mm core capture. I could use python package from photometric to capture as well. Its untested if both forms of capturing can be done at the same time or if mm called ‘dibs’ on the camera.

F2

OB1 showed 0 flow, but 20mBar pressure. I saw a drop form through the device every 5 minutes. That is slow, but significant. I flipped the pressure shutoff valve and it forced the pressure to 0 and the flow ceased. Code needs update to reflect this. \*Note this only semi worked. Having it fully calibrate itself and dictating the flow be -3 instead of 0 seemed to drive it to virtually stop flow.

## F3

This might result from an instability somewhere. Using the F2 fix, I found in a run that 2 of the 7 cycles has issues where either flow didn’t start or couldn’t stop. Same exact code. Why different results? I contacted Elveflow to talk to a software engineer about it. For this issue, I can have an issue with the fluidic system either not stopping or starting and I can fix it by simply stopping the script and restarting it. Here is a snippet of what happened when I tried a tactic of if it failed, reset the ob1 and try again with the order.

A screenshot of a computer program

Description automatically generated

Made data logger, but has not revealed anything significant yet. I was informed that usb hubs and shared power circuits between pumps and the OB1 are known to cause instability issues. I do not have a joint circuit, but I all of my devices are run from the same USB controller on the computer. Maybe that is like the external USB hub? Hard to say. The software engineer from Elveflow has been pretty unresponsive though.

I heard back from Elveflow and they and they had me swap two functions. I had start remote loop and then start PID. It needs to be reversed. I tested and it worked. Also, I tried out the data logger with the fluidics for a real run and it failed twice. Originally, I felt that the read off was accurate, but now I am not sure. On one failure, it said my flow was -3.03 instead of 500. However, it emptied the vial, which points to that it actually did the command. In addition, both of the times that it failed, the flow read off gave back a value of -3.03.

As of now (20-2-24) I believe the following to be true:

1. The PID loop is what messes up
2. The PID doesn’t mess up in vacuum and other microscope actions must be taking place too.