## Slide 1: 20-11-23 notes

## Slide 2: Focus
Only muc 2 actually stained so the rest are just evaluating focus on the autofluorescence in the sample
What slice is DAPi focused at and does it vary? Slice 6 and no it does not vary.
What slices are the rest of the colors focused at? 488 = slice 6, consistent. A555 = slice 6, consistent, A647 = slice 7, consistent, but maybe closer to slice 8.
Any drift in DAPI? Cycle 1 = slice 6, cycle 3 = slice 6 so minimal drift
Does slice change between cy1 and 3 for 488? Yes. Cy1 = slice 6, cy3 = slice 7

## Slide 3: Illumination
Very even and difficult to discern illumination alterations at tile borders.

## Slide 4: Leaks
No leaks detected. On next day pumped PBS through chamber and flowed through nicely.

## Slide 5: Needed Changes
Why aren’t things staining? Running manual stain now to see if takes into slide from the same batch
Make DAPI slice go to middle of stack. Why at slice 6 of 7? Should have been at 5.