## Slide 1: Chamber 5.0

## Slide 2: Desired Features
Flat surface for slide to rest on
Magnetic pressure plate
Access 100% of area in view window
Under-Slide magnets
Fill-Drain port
As thin as possible heater
Deep reservoir
All magnets encased (ie nothing can corrode/rust)

## Slide 3: Materials
PTC heater flexible heater
Plotting epoxy
PLA 3D printed chamber
Stainless luer lock fill-drain port
Neodymium magnets
Maxiblack adhesive film

## Slide 4: Encasement testing
Liquid UV resin can be cured strongly to UV resin fully cured items.
Will UV resin cure to PLA strongly?
Is strength enough for encasement?
Appears to give enough for encasement, but lacks bonding strength for anything structural

## Slide 5: Heating calc
Heater
Potting Epoxy
Slide
Heater output = 15W
Heater area= potting area = slide area = 1875mm^2
Potting thickness = 0.5mm
Potting Hc = 1419 J/(Kg*K)
Potting thermal conduct = 1.2 W/(m*K)
Potting density = 1.1 g/mL
Potting volume = 938 mm^3
Potting Mass = 1.1g about
Thermal conduct with dim, thermal c = (4.5*delta K) W
Potting stored heat energy at 60C = 55J
Glass slide mass = 4.7g
Heat capacity glass slide = 0.87 J/(g*K)
Stored heat at 60C glass = 143J
Combined stored heat at 60C = about 200J
Around 13 seconds to heat up